## Slide 1: Optics - Filter Properties
When using laser light sources, filters must have very sharp cutons and cutoffs since there will be many orders of magnitude more scattered laser light than fluorescence
Can specify wavelengths that filter must reject to certain tolerance (e.g., reject 488 nm light at 10-6 level: only 0.0001% of incident light at 488 nm gets through)
[RFM]
www.sliderbase.com
Notes: Optics - Filter Properties
When using laser light sources, filters must have very sharp cutons and cutoffs since there will be many orders of magnitude more scattered laser light than fluorescence
Can specify wavelengths that filter must reject to certain tolerance (e.g., reject 488 nm light at 10-6 level: only 0.0001% of incident light at 488 nm gets through)
[RFM]
www.sliderbase.com
Optics - Filter Properties
When using laser light sources, filters must have very sharp cutons and cutoffs since there will be many orders of magnitude more scattered laser light than fluorescence
Can specify wavelengths that filter must reject to certain tolerance (e.g., reject 488 nm light at 10-6 level: only 0.0001% of incident light at 488 nm gets through)
[RFM]
www.sliderbase.com

## Slide 2: Lecture Goals
This lecture is intended to describe the nature and function of optical systems
It will describe how filters operate
When filters should be used
What problems and issues must be taken into consideration
www.sliderbase.com
Notes: Lecture Goals
This lecture is intended to describe the nature and function of optical systems
It will describe how filters operate
When filters should be used
What problems and issues must be taken into consideration
www.sliderbase.com
Lecture Goals
This lecture is intended to describe the nature and function of optical systems
It will describe how filters operate
When filters should be used
What problems and issues must be taken into consideration
www.sliderbase.com

## Slide 3: Optics - Filter Properties
Long pass filters transmit wavelengths above a cut-on wavelength
Short pass filters transmit wavelengths below a cut-off wavelength
Band pass filters transmit wavelengths in a narrow range around a specified wavelength
  Band width can be specified
Neutral Density filter is a nondiscriminant intensity reducing filter
Absorption Filter is colored glass that absorbs unwanted light
www.sliderbase.com
Notes: Optics - Filter Properties
Long pass filters transmit wavelengths above a cut-on wavelength
Short pass filters transmit wavelengths below a cut-off wavelength
Band pass filters transmit wavelengths in a narrow range around a specified wavelength
Band width can be specified
Neutral Density filter is a nondiscriminant intensity reducing filter
Absorption Filter is colored glass that absorbs unwanted light
www.sliderbase.com
Optics - Filter Properties
Long pass filters transmit wavelengths above a cut-on wavelength
Short pass filters transmit wavelengths below a cut-off wavelength
Band pass filters transmit wavelengths in a narrow range around a specified wavelength
Band width can be specified
Neutral Density filter is a nondiscriminant intensity reducing filter
Absorption Filter is colored glass that absorbs unwanted light
www.sliderbase.com

## Slide 4: Optics - Filter Properties
When a filter is placed at a 45o angle to a light source, light which would have been transmitted by that filter is still transmitted but light that would have been blocked is reflected (at a 90o angle)
Used this way, a filter is called a dichroic filter or dichroic mirror
[RFM]
www.sliderbase.com
Notes: Optics - Filter Properties
When a filter is placed at a 45o angle to a light source, light which would have been transmitted by that filter is still transmitted but light that would have been blocked is reflected (at a 90o angle)
Used this way, a filter is called a dichroic filter or dichroic mirror
[RFM]
www.sliderbase.com
Optics - Filter Properties
When a filter is placed at a 45o angle to a light source, light which would have been transmitted by that filter is still transmitted but light that would have been blocked is reflected (at a 90o angle)
Used this way, a filter is called a dichroic filter or dichroic mirror
[RFM]
www.sliderbase.com

## Slide 5: Interference and Diffraction: Gratings
Diffraction essentially describes a departure from theoretical geometric optics
Thus a sharp objet casts an alternating shadow of light and dark “patterns” because of interference
Diffraction is the component that limits resolution
3rd Ed. Shapiro p 83
www.sliderbase.com
Notes: Interference and Diffraction: Gratings
Diffraction essentially describes a departure from theoretical geometric optics
Thus a sharp objet casts an alternating shadow of light and dark “patterns” because of interference
Diffraction is the component that limits resolution
3rd Ed. Shapiro p 83
www.sliderbase.com
Interference and Diffraction: Gratings
Diffraction essentially describes a departure from theoretical geometric optics
Thus a sharp objet casts an alternating shadow of light and dark “patterns” because of interference
Diffraction is the component that limits resolution
3rd Ed. Shapiro p 83
www.sliderbase.com

## Slide 6: Interference in Thin Films
Small amounts of incident light are reflected at the interface between two material of different RI
Thickness of the material will alter the constructive or destructive interference patterns - increasing or decreasing certain wavelengths
Optical filters can thus be created that “interfere” with the normal transmission of light
3rd Ed. Shapiro p 82
www.sliderbase.com
Notes: Interference in Thin Films
Small amounts of incident light are reflected at the interface between two material of different RI
Thickness of the material will alter the constructive or destructive interference patterns - increasing or decreasing certain wavelengths
Optical filters can thus be created that “interfere” with the normal transmission of light
3rd Ed. Shapiro p 82
www.sliderbase.com
Interference in Thin Films
Small amounts of incident light are reflected at the interface between two material of different RI
Thickness of the material will alter the constructive or destructive interference patterns - increasing or decreasing certain wavelengths
Optical filters can thus be created that “interfere” with the normal transmission of light
3rd Ed. Shapiro p 82
www.sliderbase.com

## Slide 7: Optical filters
Interference filters: Dichroic, Dielectric, reflective filters…….reflect the unwanted wavelengths
Absorptive filters: Colour glass filters…..absorb the unwanted wavelengths
www.sliderbase.com
Notes: Optical filters
Interference filters: Dichroic, Dielectric, reflective filters…….reflect the unwanted wavelengths
Absorptive filters: Colour glass filters…..absorb the unwanted wavelengths
www.sliderbase.com
Optical filters
Interference filters: Dichroic, Dielectric, reflective filters…….reflect the unwanted wavelengths
Absorptive filters: Colour glass filters…..absorb the unwanted wavelengths
www.sliderbase.com

## Slide 8: Interference filters
They are composed of transparent glass or quartz substrate on which multiple thin layers of dielectric material, sometimes separated by spacer layers .
Permit great selectivity.
www.sliderbase.com
Notes: Interference filters
They are composed of transparent glass or quartz substrate on which multiple thin layers of dielectric material, sometimes separated by spacer layers .
Permit great selectivity.
www.sliderbase.com
Interference filters
They are composed of transparent glass or quartz substrate on which multiple thin layers of dielectric material, sometimes separated by spacer layers .
Permit great selectivity.
www.sliderbase.com

## Slide 9
Standard Band Pass Filters
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light
www.sliderbase.com
Notes: Standard Band Pass Filters
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light
www.sliderbase.com
Standard Band Pass Filters
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light
www.sliderbase.com

## Slide 10: Standard Long Pass Filters
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters
www.sliderbase.com
Notes: Standard Long Pass Filters
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters
www.sliderbase.com
Standard Long Pass Filters
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters
www.sliderbase.com

## Slide 11: Dichroics
They used to direct light in different spectral region to different detectors.
They are interference filters , long pass or short pass.
"dichroic" Di- is Greek for two, and -chroic is Greek for color - from Greek dikhroos, bicolored
www.sliderbase.com
Notes: Dichroics
They used to direct light in different spectral region to different detectors.
They are interference filters , long pass or short pass.
"dichroic" Di- is Greek for two, and -chroic is Greek for color - from Greek dikhroos, bicolored
www.sliderbase.com
Dichroics
They used to direct light in different spectral region to different detectors.
They are interference filters , long pass or short pass.
"dichroic" Di- is Greek for two, and -chroic is Greek for color - from Greek dikhroos, bicolored
www.sliderbase.com

## Slide 12: Optical Filters
Dichroic Filter/Mirror at 45 deg
Reflected light
Transmitted Light
Light Source
www.sliderbase.com
Notes: Optical Filters
Dichroic Filter/Mirror at 45 deg
Reflected light
Transmitted Light
Light Source
www.sliderbase.com
Optical Filters
Dichroic Filter/Mirror at 45 deg
Reflected light
Transmitted Light
Light Source
www.sliderbase.com

## Slide 13: Dichroic Filters
Transmitted
Light
Reflected
Light
Filter acting as a DICHROIC
www.sliderbase.com
Notes: Dichroic Filters
Transmitted
Light
Reflected
Light
Filter acting as a DICHROIC
www.sliderbase.com
Dichroic Filters
Transmitted
Light
Reflected
Light
Filter acting as a DICHROIC
www.sliderbase.com

## Slide 14: Construction of Filters
Filter
components
Single Optical
filter
“glue”
Interference filters
www.sliderbase.com
Notes: Construction of Filters
Filter
components
Single Optical
filter
“glue”
Interference filters
www.sliderbase.com
Construction of Filters
Filter
components
Single Optical
filter
“glue”
Interference filters
www.sliderbase.com

## Slide 15: Transmission determination
Constructive and destructive interference occurs between reflections from various layers
Transmission determined by :
  thickness of the dielectric layers
  number of these layers
  angle of incidence light on the filters
www.sliderbase.com
Notes: Transmission determination
Constructive and destructive interference occurs between reflections from various layers
Transmission determined by :
thickness of the dielectric layers
number of these layers
angle of incidence light on the filters
www.sliderbase.com
Transmission determination
Constructive and destructive interference occurs between reflections from various layers
Transmission determined by :
thickness of the dielectric layers
number of these layers
angle of incidence light on the filters
www.sliderbase.com

## Slide 16: Absorptive filters
Such as coloured glass filters which absorb unwanted light.
Consist of dye molecules uniformly suspended in glass or plastic.
Remove much more of the unwanted light than do the interference filters
Will often fluoresce (not good!)
www.sliderbase.com
Notes: Absorptive filters
Such as coloured glass filters which absorb unwanted light.
Consist of dye molecules uniformly suspended in glass or plastic.
Remove much more of the unwanted light than do the interference filters
Will often fluoresce (not good!)
www.sliderbase.com
Absorptive filters
Such as coloured glass filters which absorb unwanted light.
Consist of dye molecules uniformly suspended in glass or plastic.
Remove much more of the unwanted light than do the interference filters
Will often fluoresce (not good!)
www.sliderbase.com

## Slide 17: Filters transmission
Bandpass filters: characterized by there
T max and (the Full Width at Half Maximum) FWHM
Notch filters are band pass filters in the upside down position
Long pass and Short pass filters: characterized by their T max and cuton, cutoff wavelength.
www.sliderbase.com
Notes: Filters transmission
Bandpass filters: characterized by there
T max and (the Full Width at Half Maximum) FWHM
Notch filters are band pass filters in the upside down position
Long pass and Short pass filters: characterized by their T max and cuton, cutoff wavelength.
www.sliderbase.com
Filters transmission
Bandpass filters: characterized by there
T max and (the Full Width at Half Maximum) FWHM
Notch filters are band pass filters in the upside down position
Long pass and Short pass filters: characterized by their T max and cuton, cutoff wavelength.
www.sliderbase.com

## Slide 18
Fluorescein (FITC)
400 nm
500 nm
600 nm
700 nm
Wavelength
Protein
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
www.sliderbase.com
Notes: Fluorescein (FITC)
400 nm
500 nm
600 nm
700 nm
Wavelength
Protein
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
www.sliderbase.com
Fluorescein (FITC)
400 nm
500 nm
600 nm
700 nm
Wavelength
Protein
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
www.sliderbase.com

## Slide 19: Interference filters advantages
They can be used as reflectors in two and three color analysis.
They usually do not themselves produce fluorescence.
They are available in short pass versions.
They are excellent as primary barrier filters.
www.sliderbase.com
Notes: Interference filters advantages
They can be used as reflectors in two and three color analysis.
They usually do not themselves produce fluorescence.
They are available in short pass versions.
They are excellent as primary barrier filters.
www.sliderbase.com
Interference filters advantages
They can be used as reflectors in two and three color analysis.
They usually do not themselves produce fluorescence.
They are available in short pass versions.
They are excellent as primary barrier filters.
www.sliderbase.com

## Slide 20: Interference filters: disadvantages
Lower blocking properties
Reduced passing properties
Their reflecting and passing properties are not absolute, this should be considered while dealing with multiple wavelengths
www.sliderbase.com
Notes: Interference filters: disadvantages
Lower blocking properties
Reduced passing properties
Their reflecting and passing properties are not absolute, this should be considered while dealing with multiple wavelengths
www.sliderbase.com
Interference filters: disadvantages
Lower blocking properties
Reduced passing properties
Their reflecting and passing properties are not absolute, this should be considered while dealing with multiple wavelengths
www.sliderbase.com

## Slide 21: Absorbance filters: advantages
They are inexpensive.
They have very good blocking properties.
They have very good transmission properties.
www.sliderbase.com
Notes: Absorbance filters: advantages
They are inexpensive.
They have very good blocking properties.
They have very good transmission properties.
www.sliderbase.com
Absorbance filters: advantages
They are inexpensive.
They have very good blocking properties.
They have very good transmission properties.
www.sliderbase.com

## Slide 22: Absorbance filters: disadvantages
They can only pass long wavelengths
( hence, can only block short wavelength)
Since they are made of solution of dye and glass, they can themselves produce fluorescence.
www.sliderbase.com
Notes: Absorbance filters: disadvantages
They can only pass long wavelengths
( hence, can only block short wavelength)
Since they are made of solution of dye and glass, they can themselves produce fluorescence.
www.sliderbase.com
Absorbance filters: disadvantages
They can only pass long wavelengths
( hence, can only block short wavelength)
Since they are made of solution of dye and glass, they can themselves produce fluorescence.
www.sliderbase.com

## Slide 23: Neutral density filters (N.D)
Attenuation of the light without discrimination of the wavelength.
N.D filters could be reflective or absorptive type.
They are partially silvered mirrors.
www.sliderbase.com
Notes: Neutral density filters (N.D)
Attenuation of the light without discrimination of the wavelength.
N.D filters could be reflective or absorptive type.
They are partially silvered mirrors.
www.sliderbase.com
Neutral density filters (N.D)
Attenuation of the light without discrimination of the wavelength.
N.D filters could be reflective or absorptive type.
They are partially silvered mirrors.
www.sliderbase.com

## Slide 24: Beam splitters
Absorptive N.D filters can not be used here; simply because of the heat, they will melt.
Common cover slips can be used as beamsplitters if small portion of the light is wanted, up to 5%
www.sliderbase.com
Notes: Beam splitters
Absorptive N.D filters can not be used here; simply because of the heat, they will melt.
Common cover slips can be used as beamsplitters if small portion of the light is wanted, up to 5%
www.sliderbase.com
Beam splitters
Absorptive N.D filters can not be used here; simply because of the heat, they will melt.
Common cover slips can be used as beamsplitters if small portion of the light is wanted, up to 5%
www.sliderbase.com

## Slide 25: Measuring Filter Properties
Filters must be measured at the angle they are going to be used
filters placed at 90o have different properties when they are placed at 45o
www.sliderbase.com
Notes: Measuring Filter Properties
Filters must be measured at the angle they are going to be used
filters placed at 90o have different properties when they are placed at 45o
www.sliderbase.com
Measuring Filter Properties
Filters must be measured at the angle they are going to be used
filters placed at 90o have different properties when they are placed at 45o
www.sliderbase.com

## Slide 26: Short pass and long pass filters
T
R
A
N
S
M
I
S
S
I
O
N
WAVELENGTH
SP filter
LP filter
cutoff
cuton
T max
T max
www.sliderbase.com
Notes: Short pass and long pass filters
T
R
A
N
S
M
I
S
S
I
O
N
WAVELENGTH
SP filter
LP filter
cutoff
cuton
T max
T max
www.sliderbase.com
Short pass and long pass filters
T
R
A
N
S
M
I
S
S
I
O
N
WAVELENGTH
SP filter
LP filter
cutoff
cuton
T max
T max
www.sliderbase.com

## Slide 27: Optical filter evaluation
optical filter (90o)
slit/shutter
light source
detector
monochromator
SPECTROFLUOROMETER FOR ASSESSMENT
OF OPTICAL FILTER TRANSMISSION
www.sliderbase.com
Notes: Optical filter evaluation
www.sliderbase.com
Optical filter evaluation
www.sliderbase.com

## Slide 28: Optical filter evaluation
light source
grating
beam splitter (45o)
reference PMT
slit/shutter
Optical filter (45o)
grating
Detector
PMT
www.sliderbase.com
Notes: Optical filter evaluation
www.sliderbase.com
Optical filter evaluation
www.sliderbase.com

## Slide 29: Light loss in dichroics
Reducing reliance on the in line arrangement PMTs
Placing a second fluorescence collection lens at 180o from the first one (this is more difficult in most instruments)
www.sliderbase.com
Notes: Light loss in dichroics
Reducing reliance on the in line arrangement PMTs
Placing a second fluorescence collection lens at 180o from the first one (this is more difficult in most instruments)
www.sliderbase.com
Light loss in dichroics
Reducing reliance on the in line arrangement PMTs
Placing a second fluorescence collection lens at 180o from the first one (this is more difficult in most instruments)
www.sliderbase.com

## Slide 30: Light loss by optics
The thicker the glass the less light transmitted.
Problems with glass - UV light will not pass
In UV light system use minimum optics.
www.sliderbase.com
Notes: Light loss by optics
The thicker the glass the less light transmitted.
Problems with glass - UV light will not pass
In UV light system use minimum optics.
www.sliderbase.com
Light loss by optics
The thicker the glass the less light transmitted.
Problems with glass - UV light will not pass
In UV light system use minimum optics.
www.sliderbase.com

## Slide 31: Light loss by optics
Glass can absorb UV light and can fluoresce when illuminated at that wavelength.
For excitation > 450nm, you can use glass filters, < 450nm use quartz or silica filters.
Plastic optical filters are unsatisfactory
www.sliderbase.com
Notes: Light loss by optics
Glass can absorb UV light and can fluoresce when illuminated at that wavelength.
For excitation > 450nm, you can use glass filters, < 450nm use quartz or silica filters.
Plastic optical filters are unsatisfactory
www.sliderbase.com
Light loss by optics
Glass can absorb UV light and can fluoresce when illuminated at that wavelength.
For excitation > 450nm, you can use glass filters, < 450nm use quartz or silica filters.
Plastic optical filters are unsatisfactory
www.sliderbase.com

## Slide 32: Optical filters evaluation
Use a population of appropriately stained particles and identify which filters give the maximum signal.
Spectrofluorometer amd spectrophotometers can be used as tools for assessment of optical filters.
www.sliderbase.com
Notes: Optical filters evaluation
Use a population of appropriately stained particles and identify which filters give the maximum signal.
Spectrofluorometer amd spectrophotometers can be used as tools for assessment of optical filters.
www.sliderbase.com
Optical filters evaluation
Use a population of appropriately stained particles and identify which filters give the maximum signal.
Spectrofluorometer amd spectrophotometers can be used as tools for assessment of optical filters.
www.sliderbase.com

## Slide 33: Issue to Note
Problems with filters are more likely due to using the wrong filters
Filters degrade overtime, so they have to be changed eventually
Buy high quality filters, not cheap ones
www.sliderbase.com
Notes: Issue to Note
Problems with filters are more likely due to using the wrong filters
Filters degrade overtime, so they have to be changed eventually
Buy high quality filters, not cheap ones
www.sliderbase.com
Issue to Note
Problems with filters are more likely due to using the wrong filters
Filters degrade overtime, so they have to be changed eventually
Buy high quality filters, not cheap ones
www.sliderbase.com

## Slide 34: Hints on filters
To obtain acceptable blocking of the light outside the pass band, most interference filters incorporate some absorptive elements as well as dielectric layers
www.sliderbase.com
Notes: Hints on filters
To obtain acceptable blocking of the light outside the pass band, most interference filters incorporate some absorptive elements as well as dielectric layers
www.sliderbase.com
Hints on filters
To obtain acceptable blocking of the light outside the pass band, most interference filters incorporate some absorptive elements as well as dielectric layers
www.sliderbase.com

## Slide 35: More hints...
You have to be careful while using short pass filters, specially with short wavelength, because of the transmission ability of these filters for long wavelengths (they behave like notch filters). If you have long red/near IR signals they will pass
www.sliderbase.com
Notes: More hints...
You have to be careful while using short pass filters, specially with short wavelength, because of the transmission ability of these filters for long wavelengths (they behave like notch filters). If you have long red/near IR signals they will pass
www.sliderbase.com
More hints...
You have to be careful while using short pass filters, specially with short wavelength, because of the transmission ability of these filters for long wavelengths (they behave like notch filters). If you have long red/near IR signals they will pass
www.sliderbase.com

## Slide 36: In general
Use the least number of filters necessary to reduce signal loss
Absorption result in conversion of light into heat. Thus, laser beams hitting colour glass filters may destroy these filters .
Filters have a finite lifetime.
www.sliderbase.com
Notes: In general
Use the least number of filters necessary to reduce signal loss
Absorption result in conversion of light into heat. Thus, laser beams hitting colour glass filters may destroy these filters .
Filters have a finite lifetime.
www.sliderbase.com
In general
Use the least number of filters necessary to reduce signal loss
Absorption result in conversion of light into heat. Thus, laser beams hitting colour glass filters may destroy these filters .
Filters have a finite lifetime.
www.sliderbase.com

## Slide 37: Practical consideration
In measuring weak fluorescence, we usually use field stop and interference filters behind the field stop to remove the stray light.
The shiny part (mirror side) of the filter should face the light source (collection lens)
www.sliderbase.com
Notes: Practical consideration
In measuring weak fluorescence, we usually use field stop and interference filters behind the field stop to remove the stray light.
The shiny part (mirror side) of the filter should face the light source (collection lens)
www.sliderbase.com
Practical consideration
In measuring weak fluorescence, we usually use field stop and interference filters behind the field stop to remove the stray light.
The shiny part (mirror side) of the filter should face the light source (collection lens)
www.sliderbase.com

## Slide 38: Fiber optics & optical waveguides
Fiber optics and other optical waveguides operate by total internal reflection
problems with stray light, low NA of fibers, thus low sensitivity, light collection difficult
µ
cladding
Fiber optic waveguide
www.sliderbase.com
Notes: Fiber optics & optical waveguides
Fiber optics and other optical waveguides operate by total internal reflection
problems with stray light, low NA of fibers, thus low sensitivity, light collection difficult
µ
cladding
Fiber optic waveguide
www.sliderbase.com
Fiber optics & optical waveguides
Fiber optics and other optical waveguides operate by total internal reflection
problems with stray light, low NA of fibers, thus low sensitivity, light collection difficult
µ
cladding
Fiber optic waveguide
www.sliderbase.com

## Slide 39: Optics for forward scatter
scatter
detector
iris
blocker
Laser beam
Stream in air or a round capillary
www.sliderbase.com
Notes: Optics for forward scatter
scatter
detector
iris
blocker
Laser beam
Stream in air or a round capillary
www.sliderbase.com
Optics for forward scatter
scatter
detector
iris
blocker
Laser beam
Stream in air or a round capillary
www.sliderbase.com

## Slide 40: Lecture Summary
At the conclusion of this lecture the student should understand:
Field stops and obscuration bars are necessary in systems where air or round capillaries are used
Appropriate optical filters must be placed in combinations
Filters degrade over time and should be checked
The least number of filters should be used in a system
Forward angle scatter is frequently collected using a diode detector
www.sliderbase.com
Notes: Lecture Summary
At the conclusion of this lecture the student should understand:
Field stops and obscuration bars are necessary in systems where air or round capillaries are used
Appropriate optical filters must be placed in combinations
Filters degrade over time and should be checked
The least number of filters should be used in a system
Forward angle scatter is frequently collected using a diode detector
www.sliderbase.com
Lecture Summary
At the conclusion of this lecture the student should understand:
Field stops and obscuration bars are necessary in systems where air or round capillaries are used
Appropriate optical filters must be placed in combinations
Filters degrade over time and should be checked
The least number of filters should be used in a system
Forward angle scatter is frequently collected using a diode detector
www.sliderbase.com